## Slide 1: Thesis Presentation
Using Software Defined Radios for radiometer applications
Matthew E. Nelson
Committee Members:
Phillip Jones
John Basart
Mani Mina
Spring 2014

## Slide 2: Thesis overview
Introduction
  Software Defined Radios
  GNURadio Software
  Current ISU Radiometer
  Related Works
General theory of operation
  Software Defined Radios (SDR)
  GNURadio Software
  RF Front end

## Slide 3: Thesis overview
GNURadio and N200 TPR
  Hardware
    Requirements
    Noise Temperature Considerations
    Existing ISU Radiometer RF Front end
  Software
    Requirements
    Obtaining signal information and GUI Controls
    Implantation of a TPR in software
    Data storage and Display

## Slide 4: Thesis overview
Testing and Verification
  Square-law Detector
  SDR Tests
  E E 518 Lab tests
  LN2 Tests

## Slide 5: Thesis overview
Performance and Evaluation
  Sensitivity, Accuracy and Stability
  Required Performance Requirements
  Square-law Detector Performance
  SDR Performance
Conclusion and Future work

## Slide 6: General theory
Radiometer theory
SDR Role

## Slide 7: Radiometer theory
Radiometers measure power
Specifically radiometers measure the total noise power
The total noise power measured can be converted to a noise temperature
This information can be used for remote sensing

## Slide 8: Radiometer theory
Radiometer chain
  Antenna/Source
  Filter
  Amplify
  Detect
  Measure

## Slide 9: Radiometer theory
Noise consideration
  Each component will add unwanted noise
  Steps are taken to minimize
    Low Noise Amplifiers (LNAs)
    Low noise figures on components

## Slide 10: Radiometer theory
Stability
  A radiometer must be stable
  Different methods for achieving this
    Dicke Radiometer
  ISU Radiometer
    Precise temperature control
    We assume our source is stable

## Slide 11: SDR Role
Current ISU Radiometer

## Slide 12: SDR Role
Current ISU Radiometer
  Directly digitizes the signal
  A/D Under samples the information
    Can only recover power information
    No spectrum information is preserved
  Is currently not functioning
    Issues with the FPGA

## Slide 13: SDR Role
SDR replacement
  Replaces current A/D and FPGA
  Off the shelf component
  Standard equipment
  Open source firmware and software
  Smaller package

## Slide 14: SDR Role
Amplification and filtering
Power Divider
Square-Law Detector
N200
PC
PIC32
DBSRX2
DBSRX2
A/D
FPGA
GNURadio
Python Datalogger
Data
Storage
Matlab
~ NF of 3.67 dB (~ 385 K)
This includes antenna and all filters and LNAs
Data provided by University of Michigan
Total amplification of 71.4 dB
N200 uses the DBSRX2 daughterboard
  800 MHz to 2.4 GHz
  ~ NF of 5 dB (626 K)
A/D
FGPA
1.8 Million gates
32-bit RISC processor
Square Law
-60 to 1 dBm
PIC32
Up to 80 MHz operation
10-bit A/D
I2C, SPI, 4xUARTs
512K FLASH
A/D Converter
100 Mbps 14-bit

## Slide 15: SDR Role
GNURadio implements a radiometer in software
  Digital Square-Law
  Filter and integration
This information is saved for later analysis
  Raw IQ values
  Total Power values
  Spectral information

## Slide 16: SDR Role
GNURadio allows us to control the system
  Frequency
  Gain
    Within the SDR
  Bandwidth (Sample time)
  Integration time

## Slide 17: SDR Role

## Slide 18: SDR Role
GNURadio Receives IQ data from source
  N200 SDR
  Simulated Noise Source
  Recorded IQ values

## Slide 19: SDR Role
Software square-law detector
  Squares and sums the values
  Takes in complex value
  Outputs a float value
  Replaces a typical Square-law detector

## Slide 20: SDR Role
Software Low Pass Filter
  Smooth out data
  Acts as our integrator
  Integration time sets Alpha values

## Slide 21: SDR Role
Decimation
  Reduces sample rate
  Does not affect bandwidth
  Reduces CPU requirements

## Slide 22: SDR Role
Output
  Display data in real-time (ticker tape fashion)
  Stores data in a file
  “valve” allows for turning on/off recording

## Slide 23: SDR Role
Calibration
  Pre-calculated points can be entered
  Allows data to be displayed as a noise temperature

## Slide 24: SDR Role
GUI Interface
  Allows for control of most SDR functions
  Allows for data input
  Allows for graphing and data output

## Slide 25: SDR Role
SDR Control
  Set Frequency
  Adjust gain in SDR

## Slide 26: SDR Role
SDR Control
  Sample Rate or BW
  High BW results in higher CPU and Disk I/O
  Is somewhat limited by RF Frontend

## Slide 27: SDR Role
SDR Control
  Integration time on IIR Filter
  Defaults to 2 seconds

## Slide 28: SDR Role
SDR Control
  Artificial gain
  Simple multiplication to the digital data

## Slide 29: SDR Role
SDR Control
  Calibration values
  Allows for data to be displayed as a noise temperature

## Slide 30: SDR Role
SDR Control
  Inject simulated noise
  Adjust the amplitude of the simulated source

## Slide 31: Research Infrastructure
Verify SDR results
  Square-law detector used
  These are traditionally used in analog radiometers

## Slide 32: Square-law system
Square-Law Detector
RF Signal
Square-Law Detector
PIC32
RF Signal

## Slide 33: Square-law system
Sanity check
  Traditional radiometer
Analog signal feed to 10-bit A/D on PIC32
  10-bit resolution means we have a resolution of ~ .3 K
    Assuming a 0-313 K range of temperatures
Data is sent from PIC32 to Host PC
  Serial data interface
  Currently uses an xBee to wireless transmit
Python script records data and timestamps it

## Slide 34: Infrastructure Testing
Two files are saved to disk
  GNURadio data (Binary INT16 File)
  Square-law from PIC32 (ASCII/CSV File)
ASCII/CSV file is human readable
GNURadio file is not, requires program
A Matlab script can read both

## Slide 35: Infrastructure Testing
Sample test
Signal Generator
N200
Square Law
.dat file
.csv file
Matlab Script

## Slide 36: Infrastructure Testing
Sample test
  Signal generator at 850 MHz
    Due to power divider used
  Power divider splits signal
    Approx. 6-7 dB loss
      Due to outside of divider range (200-500 MHz)
  Signal recorded on both

## Slide 37: Infrastructure Testing
Results – Youtube videos

## Slide 38: Infrastructure Testing

## Slide 39: Infrastructure Testing
Results
SDR has a slighter higher sensitivity than the square law
The system is able to record both sets of data and using a Matlab script we can analyze both sets of data

## Slide 40: LN2 testing
Liquid Nitrogen (LN2) is used to test the ability of the radiometer to detect a wide range of noise temperatures
A 50-ohm load is submerged into a LN2 bath
The results are then recorded by the SDR
This also allows us to calibrate the radiometer with known temperatures

## Slide 41: LN2 testing

## Slide 42: LN2 testing
Once two points were obtained, we solved for the two unknowns
These calibration points can then be entered into GNURadio
GNURadio can now output data as a noise temperature in K

## Slide 43: LN2 testing

## Slide 44: Theory on operation
In a Software Defined Radio, we represent the signal by the amplitude (I) and the phase (Q) of the signal.
  By combining these values, we can recreate the signal.
The MAX2112 directly converts the RF signal to analog I and Q values
This is then sampled and sent to GNURadio

## Slide 45: Theory on operation
In GNURadio we use the Complex^2 block
The Complex^2 block does the following
This effectively takes the signal and squares it to find the amplitude of the signal

## Slide 46: Theory on operation
As mentioned before, the raw data from the square of the signal will be erratic
We want to filter and smooth this data
We also want to integrate the signal as well
A very common filter in the analog world is a RC filter, which also acts as an integrator

## Slide 47: Theory on operation
But how do we do an RC filter in the digital domain?
We use an IIR Filter

## Slide 48: Theory on operation
To determine Alpha in GNURadio, we request the integration time from the user and the sample rate (BW)

## Slide 49: Theory on operation
The base requirement for the radiometer is a sensitivity (or ) of 1 K. To achieve this we can adjust our integration time.

## Slide 50: Theory on operation
The ISU Radiometer front = 385 K
N200 front end = 626 K
= 926 K
Typical = 300 K
BW to 20 MHz (current radiometer setup)
If we want a of 1 K, then solving for gives us 0.086 seconds for the integration time

## Slide 51: Theory on operation
If we set to 1 second, the bandwidth to 20 MHz, our is .29 K, which is quite acceptable.

## Slide 52: EE 518 Lab
Experiment
Radiometer rotated up to the sky
Rotated at set angels and data recorded
Repeated and changed states (Ref. load, noise diode, Antenna)
Also did a continuous sweep from 180 to 0
Day was a clear sunny day
Plate temperature was constant 293 K

## Slide 53: EE 518 Lab

## Slide 54: EE 518 Lab
The reference load should be consistent, instead it varied as the radiometer moved
Reference load and antenna data was almost identical
This leads us to believe the switch was not working correctly
There appears to be a change in gain as the radiometer moves
The temperature increased with the noise diode turned on, as expected
Spike in the data however is disturbing

## Slide 55: Future Work
Tweaks to GNURadio interface
  Refinements to control
  Self-calibration
Integration of square-law detector
  Inject data into GNURadio
Integration to other radiometer systems
  Temperature control
  Switching noise diode

## Slide 56: Future Work
Add correlation
  Further testing with two DBSRX2 modules

## Slide 57: Radiometer Refit
RF Front End needs re-work
  New LNAs
  Fix interference issue
Power Sub-System
  Adjust for new LNAs
  Adjust for new digital HW
New microcontroller
  PIC32 integration
    Control TECs
    Read in temperature sensors

## Slide 58: Extras
Extra slides

## Slide 59: Theory on operation
In the power series, we can replace “e” with
We can now use the following trig identities

## Slide 60: Theory on operation
Substitute the trig identities

## Slide 61: Theory on operation
We are only interested in the voltage across the load resistor
Substitute
By passing this through a simple low pass filter, we can output the DC output voltage

## Slide 62: Theory on operation
Simplify the constants to
Note that a DC term has appeared

## Slide 63: Research Infrastructure

## Slide 64: Theory on operation
Substitute Q into the equation
now substitute V

## Slide 65: Theory on operation
Since we are in the digital domain, it needs to be discrete which we will represent as as the samples of and
Now substitute

## Slide 66: Theory on operation
Now rearrange which gives us the recurrence relation

## Slide 67: Software Defined Radio
N200 Specs
Ettus USRP N200 Unit
  Up to 50 MHz BW on RX
1 Gbps interface
DBSRX2 Daughter board
  800 MHz to 2350 MHz receive
  We listen at 1.4 GHz
  Noise Figure of 5 dB

## Slide 68: Software Defined Radio

## Slide 69: Square Law detector
Analog Devices ADL5902 Square-law detector

## Slide 70: Square Law detector
Tested linearity of the square law detector

## Slide 71: Square Law detector
Analog Devices ADL5902 Square-law detector

## Slide 72: EE 518 Lab
The first “real” test of the square-law detector
ADL5902 measured RF power
Analog signal sent to PIC32 A/D converter
  10-bit resolution
  Sample rate ~ 100 kHz
PIC32 processed and sent A/D values
  Used xBee to send to laptop

## Slide 73: Research Infrastructure
Current RF Front End
~ NF of 3.67 dB (~ 385 K)
  This includes antenna and all filters and LNAs
  Data provided by University of Michigan
Total amplification of 71.4 dB
Also note, the Power Divider will add another ~21 K to the system

## Slide 74: EE 518 Lab

## Slide 75: EE 518 Lab

## Slide 76: EE 518 Lab

## Slide 77: EE 518 Lab
Results – N200

## Slide 78: EE 518 Lab
Results – N200
As seen in the video, there is a problem with the RF Front end
It seems to change as the physical radiometer rotates
In addition it is clear that something is creating a signal in the RF chain
  We currently think this is due to a faulty LNA(s)
This…is not a good thing
Further investigation is required, but is outside the scope of this thesis

## Slide 79: Research Infrastructure
RF Signal fed into DBSRX2 daughter cards
  Up to 2 may be installed

## Slide 80: Research Infrastructure
MGA6x564
PGA (Programmable Gain Amplifier)
.9 dB NF
Up to 22 dB Gain

## Slide 81: Research Infrastructure
MAX2112
  Direct conversion IC
    Downconverts directly to baseband IQ
  Outputs analog I and Q values
  Adds an additional ~4 dB Noise Factor

## Slide 82: Research Infrastructure
Time stamping

## Slide 83: Square-law system
  Integration time is controlled by setting the clock frequency for the A/D
  so for example if we want 5 ms
  PIC32 currently sends data via RS232
    Other methods could be used (Ethernet)

## Slide 84: Theory on operation
A square-law detector is usually a diode
  In software, we simply model this as a square function, why?

## Slide 85: Theory on operation
We can show that this represents a DC term plus all the harmonics. If we pass this through a low pass filter, we can output only the DC output
Therefore, this is a DC voltage proportional to the square of the amplitude of the input RF signal.

## Slide 86: Theory on operation – X^2
We assume the following
  A small signal is operating in the knee of the diode’s V-I Curve
  The V-I curve of the series combination of the Diode and the load resistor can be represented by a power series of the following form

## Slide 87: Theory on operation
With a digital filter, we move to the time domain
Have represent the charge on the capacitor at time t

## Slide 88: Theory on operation
This is saying that the discrete-time implementation of a simple RC low-pass filter is the exponentially weighted moving average

## Slide 89: Theory on operation
We can then take and set that to a constant, called alpha (

## Slide 90: Theory on operation
This alpha then gives us our equivalent time constant (RC value) that we would expect in an analog filter. In the digital domain, alpha is referred to as the smoothing factor and is

## Slide 91: EE 518 Lab
Results – N200
After the initial lab, the N200 was hooked up to record a sample
  20 MHz sample rate
  Stored to laptop HDD
  Viewed later after sample recorded

## Slide 92: EE 518 Lab
Results – N200

## Slide 93: EE 518 Lab
Conclusion
Lab was a “success” in that data was recorded
Our original assumption that the RF front end was “ok” appears to be incorrect
While the data was inconclusive and not what we expected, it validated using a SDR to do further analysis on the signal

## Slide 94: EE 518 Lab
Results
Good news
  Data was recorded and has some consistency
  Noise diode is clearly working
Bad news
  Does not appear we are switching sources
  Odd spike in Noise Diode data
  Something is not quite right with the RF Front end

## Slide 95: EE 518 Lab
Results
Data was inconclusive
  One hypothesis is the RF switch did not function as intended
Data did not “match up” to what was expected
Changes however were measured
  High confidence that square-law detector and PIC32 did their job
Data recorded with the SDR shed some light on some issues

## Slide 96: Infrastructure Testing
Simulink file
  Created to help in determine system noise temp.
Simulink has this built into SimRF
  Unfortunately ISU does not have a license for this
Created simple “m-blocks”